Comparison of T° measurements @ Humain
Thomas Goossens (CRA-W) - t.goossens@cra.wallonie.be
16 May, 2018
# Context & objectives
## Integration of the RMI + Pameseb AWS networks
- Agromet : providing hourly gridded (1km²) weather datasets for agricultural decision support system
- Possibility of RMI stations integration for better spatialization ==> interoperability ?
## Assessing the agreement level between the 2 types of AWS
- previous work by Geoff Jenkins described a comparison method based on regression model correlation study.
- as demonstrated by Giavarina, correlation studies is not appropriate to study the differences.
- The alternative Bland-Altman analysis proposed by the authors is used in this study.
# Data presentation
## Period of interest
2 years of hourly records at Humain from 01-nov-2015 to 01-nov-2017
## Statistics :
id	sid	poste	sum	mean	min	max	sd	var	q25	q50	q75
31	61	Humain	176438.2	10.18	-9.80	33.7	7.14	50.99	4.80	9.90	15.20
64	1000	Humain	175695.3	10.14	-9.62	33.2	6.97	48.52	4.94	9.94	15.02

tsa summary stats
## Time series - full

Time series
## Time series - january 2016

Time series - january 2016 zoom
## Time series - august 2016

Time series - august 2016 zoom
## Density

Density
## Scatter

Scatter plot
# First insights
## Similarity
Datasets seem in accordance ...
- but to which extent ?
- How can we quantify this degree of similarity ?
- Even if a certain degree of similarity : Pameseb is globally warmer
- For the daily max _above 10 °C : Pameseb warmer than RMI and the opposite below 10°C
- The disparities are generally found around the hours corresponding to daily max temperature.
# Investigating the temperature difference
## Bland-Altman analysis - concept
We can use a Bland Altman analysis to measure the agreement level between the 2 stations
It consists to take observations pair by pair and to plot their means vs. their differences.
In our case we choose to plot IRM - Pameseb.
## Bland-Altman Plot

## BA interpretation
- For most individuals, measurements made by 2 stations are more likely to be far apart between [-0.78 ; 0.69] °C.
- 0 is not in the 95% CI of the mean difference [-0.77 ; -0.037 ] ==> It exists a mean negative bias of -0.043
- The scatter presents a negative trend going from positive bias to negative bias ==> proportional bias. The methods do not agree equally through the range of measurements. The highest absolute differences are observed at the highest mean temperatures.
- Below 10 °C, IRM tends to be warmer (T° difference fluctuates above 0°C) and above 10°C, Pameseb tends to be warmer (T°difference fluctuates below 0°C)
## A potential explanation : station design
The 2 stations are equipped with different temperature sensors and are actually different in their design. While the RMI is equipped with a Stevenson radiation screeen mechanically ventiled, Pameseb is equipped with a custom screen without mechanical ventilation.

# Potential effect of the mechanical ventilation
## Supposition
The lack of ventilation system at Pameseb does not allow to sufficiently compensate the effect of direct solar radiation, potentially leading to an overheating of the Pameseb shelter compared to RMI during situations of high irradiance and low wind.
If this supposition is true and is the only explanation ==> the negative differences should be inexistant or at least less pronounced at night
## Bland-Altman Plot : night only situations

## BA night only situations - interpretation
- For most individuals, measurements made by 2 stations are more likely to be far apart in a narrower range comprised between [-0.35 ; 0.59] °C.
- 0 is not in the 95% CI of the mean difference [-0.34 ; 0.12 ] ==> It exists a mean positive bias of 0.12 (IRM warmer than Pameseb this time)
- The scatter still presents a trend going from positive bias to negative bias.
- ==> The ventilation system is not the only explanation to the proportional bias. The sensitivity of the sensors themselves and/or non-linear differences in radiation transfer effects from the shelters towards the sensors could also possibly explain the observed differences.
- ==> The ventilation system could however play a significant role during the time of the day were the radiation is at its maximum. What happens around daily max temperature values ?
## Bland-Altman Plot : daily max +/- 2h situations

## BA daily max +/- 2h situations - interpretation
- For most individuals, measurements made by 2 stations are more likely to be far apart in a range comprised between [-1.1 ; 0.48] °C.
- 0 is not in the 95% CI of the mean difference [-1 ; -0.28 ] ==> It exists a mean negative bias of -0.29 (Pameseb globally warmer)
- The scatter still presents a negative trend but no more positive bias (Pameseb is always warmer).
- ==> The ventilation system seems to effectively play a role when the received radiation is at its maximum.
## Temperature difference Summary
- The lack of ventilation system at Pameseb is not the only explanation to the BA plot pattern observed when taking all the records in account.
- It is however able to explain what is going on around the daily_max temperature values
- But we need to invoke a second physical process able to explain the mean positive bias for the temperature below 10°C : radiative effect of the wood shelter ? Warmer air from below blown in the shelter ?
# Solving the disagreement
## Idea
- Applying a temperature correction model to Pameseb stations prior to the integration of the 2 networks in the spatialization process (we consider RMI as the gold-standard)
- The value of the correction to apply to each observation will be predicted by a correction model able to predict the observed temperature difference by learning from any combination of wind speed, irradiance and temperature measured at Pameseb.
- corr_pameseb_temp = RMI_temp + predicted_temp_diff
## Correction strategies
Regarding what we have learned from the Bland-Altman plots, we can use 2 correction strategies :
- 1st focusing on the daily temperature peaks differences : only corrects observations corresponding to [-2h; daily max ; +2h] ==> train set = observations corresponding to High clearness index values (ci > q70(ci)) and test set = observations corresponding to [-2h; daily max ; +2h]
- 2nd non-specific approach of the temperature differences. ==> train and test sets = random split ("holdout") of the whole dataset.
## Predictive performance of the models - daily temperature peaks

## Predictive performance of the models - non-specific

## Predictive performance of the models - interpretation
- In both approaches, the model that presents the lowest mean square error is the one using the 3 variables (wind speed, irradiance and temperature) as predictors.
- For the non-specific approach, the learner that performs the best is the linear regression with a mse of 0.0900426
- For the daily temperature peaks approach, the learner that performs the best is the Fast k-Nearest Neighbor with a mse of 0.1030019
- ==> What is the best correction approach ? Let's compare the predictions outputs
## Corrected Bland-Altman Plot - non-specific

Bland-Altman plot - Pameseb corrected by non-specific approach
## Corrected BA daily - Interpretation
- For most individuals, measurements made by 2 stations are more likely to be far apart in a narrower range comprised between [-0.6 ; 0.6] °C.
- 0 is in the 95% CI of the mean difference [-0.59 ; 0.0066 ] ==> no more bias !
- No more negative trend is observable
## Corrected Bland-Altman Plot - daily temperature peaks

Bland-Altman plot - Pameseb corrected by daily temperature peaks approach
## Corrected BA non-specific - Interpretation
- For most individuals, measurements made by 2 stations are more likely to be far apart in a narrower range comprised between [-0.67 ; 0.67] °C.
- 0 is in the 95% CI of the mean difference [-0.66 ; 0.0062 ] ==> no more bias !
- But the negative trend is still visible...
## Corrected Time-series Plot - august - daily temperature peaks

Time series - august 2016 zoom Pameseb corrected by daily temperature peaks approach
## Corrected Time-series Plot - august - non-specific

Time series - august 2016 zoom Pameseb corrected by non-specific approach
## Corrected Time-series Plot - january - daily temperature peaks

Time series - january 2016 zoom Pameseb corrected by daily temperature peaks approach
## Corrected Time-series Plot - january - non-specific

Time series - january 2016 zoom Pameseb corrected by non-specific approach
# Conclusions, perspectives and colofon
## Conclusions
- Using Bland-Altman to assess the difference between two datasets, we were able to demonstrate that 2 stations does not record the temperature in similar ways
- Using the power of mlr, we are able to quickly test various correction models and compare their predictions
- Using Pameseb temperature, wind speed and irradiance as explanatory variables, we are able to predict the temperature difference with a mse < 0.10 and to suppress the observed proportional bias
## Perspectives
- improvement 1 : develop 2 correction models ==> one below 10°C and another one above 10°C ?
- Further work : can this correction process be applied asis to all our stations ? For this purpose, we can could compare irradiance patterns from close stations : (Ernage vs. Sombreffe) + (Buzenol vs. Ruette). if they behave the same as Humain ==> correction model transposable
- Once finished, we plan to submit this work to the "Weather" Journal of the RMETS : Quantitative intercomparison of 2 years hourly records of temperature and relative humidity measured by 2 different types of professional-grade automatic weather stations at Humain, Belgium
## Colofon
- This report is still in major revision phase. Consider it as a first draft version.
- This document was generated using R software with the knitr library
- The interactive plots are rendered from ggplot by plotly.
- The model definition was performed using the mlr package
- The source code of the analysis and this presentation is availbale on github
## Terms of service
To use the AGROMET API you need to provide your own user token.
The present script is available under the GNU-GPL V3 license and comes with ABSOLUTELY NO WARRANTY.
Copyright : Thomas Goossens - t.goossens@cra.wallonie.be 2018.
(This document was generated using R software with the knitr library).